## Slide 1: GaAs X-ray Detectors: properties compared against Si and CdTe
Stephen Shefsky
25 June 2014

## Slide 2: Semiconductor Properties

## Slide 3: Detector Dimensions
 |  | Si | GaAs | CdTe | notes | ref | 
area | mm^2 |  | 0.04 |  |  | Eggert | 
 |  |  | 0.79 |  |  | Lowe | 
 |  |  | 1.77 |  |  | Owens | 
 |  |  | 25.00 |  | epi-GaAs | Lowe | 
 |  | 25 |  | 25 |  | Amptek | 
thickness | um |  | 60-100 |  | CVD | Eggert | 
 |  |  | 40-400 |  | epi | Lowe | 
 |  |  | 250 |  | gas phase epi | Fedorov | 
 |  |  | 200 |  | HVPE | Zhilyaev | 
 |  |  | 40 |  |  | Lowe | 
 |  |  | 40 |  |  | Owens | 
 |  | 500 |  | 1000 |  | Amptek | 

## Slide 4: Performance: resolution
 |  | Si | GaAs | CdTe | notes | ref
resolution @ 5.9 KeV | eV |  | 220 |  | 0.04 mm^2, -30 C | Eggert
 |  |  | 200 |  |  | Fedorov
 |  |  | 219 |  | -31 C | Lowe
 |  |  | 630 |  | -40 C | Owens
 |  |  | 728 |  | -35 C | Owens
 |  |  | 266 |  | room temp | Lowe
 |  |  |  | 240, 249 | -40 C | Lowe
 |  | 125-140 |  |  | SDD | Amptek
 |  | 145-230 |  | 290 | PIN, CdTe | Amptek
resolution @ 60 KeV | eV |  | 487 |  | room temp | Lowe
 |  |  | 600 |  |  | Fedorov
 |  |  |  | 600 | cooled | Lowe
 |  |  |  | 600 |  | Amptek

## Slide 5: Performance: other
 |  | Si | GaAs | CdTe | notes | ref
leakage | nA/cm^2 |  | <4 |  | 100 V, room temp | Lowe
 |  |  | 2.3 |  | -40 V, -35 C, | Owens
 |  |  | <100 |  |  | Zhilyaev
dead layer thickness | um |  | 5 |  |  | Owens
residual impurity conc | cm^-3 |  | 1e12 |  |  | Fedorov
 |  |  | 1e12 |  |  | Zhilyaev

## Slide 6: X-ray stopping power
With a higher density and a much higher average atomic number GaAs is much more efficient at stopping x-rays than Si.
The x-ray stopping power of GaAs is intermediate between those of Si and CdTe.
As a rough rule of thumb, 20 µm of GaAs is equivalent in x-ray stopping power to 500 µm of Si.

## Slide 7: Depletion Layer Thickness
w increases with Vb until the detector’s thickness is fully depleted. The depletion layer determines the detector’s active (sensitive) volume. To achieve full depletion we need adequate resistivity ρ and bias Vb.

## Slide 8: Capacitance
where
ε0 = 8.854e-2 pF/cm and (ε/ ε0) is 11.9 for Si and 12.9 for GaAs.
With increasing bias voltage w increases so C decreases (good!) until full depletion is reached. Lower detector capacitance means faster preamp operation, reduced noise, and better resolution.
A simple capacitance measurement may be used to determine depletion depth through the formula above.

## Slide 9: Connection of resistivity to carrier concentration
Higher resistivity material results in low carrier concentration, increased depletion layer thickness, and lower (bulk) leakage current.

## Slide 10: Reverse Current
total reverse current
= diffusion
+ surface leakage
+ generation leakage
Under typical usage conditions reverse current is dominated by generation leakage from the detector’s active volume. However, surface leakage has a stronger temperature coefficient than generation leakage, so surface leakage becomes more significant as detector temperature rises.

## Slide 11
Using the formula connecting carrier concentration n with resistivity ρ:
Generation (bulk) leakage
Note the bulk leakage’s direct dependence on the detector’s active volume Aw and its inverse dependence on the material’s mobility-lifetime product µeτ.

## Slide 12
Intrinsic carrier concentration
Nc and Nv are the effective density of states for the conduction and valence bands.
A more readily useable approximation:

## Slide 13
Intrinsic carrier concentration, continued
Since the band gap energy is in the exponent it has a profound influence on the temperature dependence of the intrinsic carrier concentration. The larger band gap of GaAs relative to Si (1.42 eV versus 1.12 eV) means that GaAs should have a lower intrinsic carrier concentration at any given temperature, and a less steep increase of carrier concentration as the temperature increases. Lower carrier concentration is generally associated with higher resistivity.

## Slide 14
Intrinsic carrier concentration, continued
Estimating that c ~ 5e16, this formula shows that the intrinsic carrier concentration for Si (Eg = 1.12 eV) at -30 C is roughly 5e-8. Using the same formula for GaAs (Eg = 1.42 eV), we find that the intrinsic carrier concentration of GaAs is about 3 orders of magnitude lower than Si at the same temperature.
Interestingly, the formula predicts that the intrinsic carrier concentration for GaAs at 300 K (~ room temperature) is as low as that of Si at -30 C.
Intrinsic carrier concentration is the theoretical limit to which the actual carrier concentration may approach but cannot beat.

## Slide 15
Electron drift length
Note the direct influence of the mobility-lifetime product on drift length. For charge collection to be “complete” the drift length λe must be much longer than the longest possible electron drift path from the point of initial x-ray interaction to the collection anode.
If charge collection is incomplete we expect the energy dispersive x-ray peaks to become shifted to lower energy, broadened, and asymmetric.

## Slide 16
Electron drift length: examples
Example 1: GaAs PIN detector, 100 um depletion thickness, 100 V bias.
E = (100 V)/(100 um) = 1e4 V/cm
λe = (1e-4 cm2/V)*(1e4 V/cm) = 1 cm.
Since the electron drift length is about 100 times the longest electron path we can expect the (electron) charge collection to be quite good.

## Slide 17
Electron drift length: examples
Example 2: GaAs drift detector, 3 mm radius, 300 V bias.
E = (300 V)/(3 mm) = 100 V/mm = 1000 V/cm (on average)
λe = (1e-4 cm2/V)*(1000 V/cm) = 0.1 cm = 1 mm
Since electron paths (which range from 0 to ~3mm) are typically longer than the electron drift length, we should expect charge collection to be poor.

## Slide 18
Escape Peaks
Escape peaks result from the escape of fluorescence x-rays produced within the detector’s sensitive volume. Any photon that strikes the detector with an energy greater than one of the detector’s K absorption edges (for GaAs these are Ga-K = 10.37 and As-K = 11.86) will produce escape peaks at the original photon energy minus each of the detector’s K x-ray energies, the most prominent in this case being the Ga-Ka (9.25) and As-Ka (10.24).

## Slide 19
Escape Peaks
Example 1: The spectrum below comes from a GaAs x-ray detector. To the left of the Am-241 gamma ray (59.54) we see significant escape peaks at their expected positions (59.54 – 9.25, 59.54 – 10.24, etc.).

## Slide 20
Escape Peaks
Example 2: Mo-K x-rays from stainless steel impinge on a GaAs detector.
We will see not only the Mo-Ka and Mo-Kb peaks (17.48, 19.61 KeV), but also escape peaks at (Mo-Ka – Ga-Ka) = 8.23, (Mo-Ka – Ga-Kb) = 7.22, and so forth.
A table of escape peak energies that result from this very simple example appears below. Consider the potential interference of these escape peaks with the important Mn-Ka (5.89), Co-Ka (6.93), Ni-Ka (7.48), Cu-Ka (8.06), and Pb-La (10.5) lines.